SCOPE 1 ЛИК

Spectri Ltd.

SPECTRI LAB Testing and calibration laboratory

Management address:

1421 Sofia, Lozenets, 30 Sv. Teodosii Tarnovski Str, floor 1, apt.3

Laboratory address:

1700 Sofia, Studentski Komplex, 5A Baku Str, office 5-2,3

To perform testing of:

| Type of the scope: flexible |  |  |  |
| --- | --- | --- | --- |
| № | Tested products | Type of test / characteristic | Test methods (standard/ validated method) |
| 1 | 2 | 3 | 4 |
| 1. | Environmental Noise* | Equivalent noise level (Equivalent sound pressure level), dBA | БДС ISO 1996-2; БДС 15471; БДС ISO 8297 Methodology for determining the total sound power emitted by an industrial plant in the environment and determining the noise level at the site of impact |
|  |  | Maximum A-weighted noise level (Maximum A-weighted sound pressure level), dBA | БДС 15471; БДС ISO 1996-2; |
|  |  | Noise level at the place of impact, dBA | Methodology for determining the total sound power emitted by an industrial plant in the environment and determining the noise level at the site of impact |
|  |  | Total sound power level, dBA | БДС ISO 8297 Methodology for determining the total sound power emitted by an industrial plant in the environment and determining the noise level at the site of impact |

Flexible scope: Implementing a new version of standards/documents or standards/ documents replacing them is allowed. An updated list of standards/documents and their dated versions is provided by CAB.

* The testing is performed at a location requested by the client.

Reference:

Methodology for determining the total sound power emitted by an industrial plant in the environment and determining the noise level at the site of impact, approved by Order № РД-613/08.08.2012 of Minister of Environment and Water.

To perform calibrating of:

| Type of the scope: fixed |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| № | Measuring instrument | Measured value, unit of measure | Measurement range | Measurement uncertainty | Calibration method |
| 1 | 2 | 3 | 4 | 5 | 6 |
| 1. | Acoustic (sound) calibrators | sound pressure level, dB for nominal frequencies 250 Hz and 1 kHz | 94 dB, 114 dB, 124 dB (relative to 20 μPa) | 0.15 dB | ILM 01 (БДС EN 60942) |
| 2. | Acoustic (sound) multifunction al calibrators | sound pressure level, dB for nominal frequencies from 31.5 Hz to 16 kHz | 94 dB, 104 dB, 114 dB (relative to 20 μPa) from 31.5 Hz to 4 kHz from 8 kHz to 16 kHz | 0.2 dB 0 4 dB | ILM 01 (БДС EN 60942) |
| 3. | Sound level meters | sound pressure level, dB for nominal frequencies 31.5 Hz to 12.5 kHz (acoustic calibration) | from 94 dB to 114 dB (relative to 20 μPa) from 31.5 Hz to 2 kHz From 4 KHz to 12.5 kHz | 0.3 dB 0.6 dB | ILM 02 (БДС ЕN 60651; БДС EN 60804; БДС EN 61672-1; БДС EN 61672-2; БДС EN 61672-3) |
|  |  | sound pressure level, dB for nominal frequencies 20 Hz to 16 kHz (electronic calibration) | from 20 dB to 140 dB (relative to 20 μPa) | 0.4 dB |  |
| 4. | Personal noise dose meters | sound pressure level, dB for nominal frequencies: from 63 Hz to 8 kHz (acoustic calibration) | from 60 dB to 120 dB (relative 20 μPa) - from 63 Hz to 2 kHz - for 4 kHz - for 8 kHz | 0.4 dB 0.8 dB 1, 3 dB | ILM 03 (БДС EN 61252) |
|  |  | Sound pressure level, dB for nominal frequencies from 63 Hz to 8 dB (electronic calibration) | from 60 dB to 140 dB | 0.4 dB |  |

Note:

The calibration of the measuring instruments shall be done only in the laboratory.

References:

ILM 01 - Internal laboratory calibration method for sound calibrators (БДС EN 60942)

ILM 02 - Internal laboratory calibration method of sound level meters (БДС EN 60651, БДС EN 60804, БДС EN 61672-1, БДС EN 61672-2, БДС EN 61672-3)

ILM 03 - Internal laboratory methodology for calibration of personal noise dose meters (БДС EN 61252).